## Slide 1: ME 366 Lungs Performance Analysis Using Spirometry
Submitted By: 1710007 1710010 1710023
Bangladesh University of Engineering and Technology
A18

## Slide 2: Spirometry & Uses
Spirometry is a test to measure the volume of exhaled air
The Value that we are concerned in this project is "Forced Vital Capacity"
Forced vital capacity (FVC) is the maximum amount of air forcefully exhaled from the lungs
Forced Vital Capacity (FVC) depends on age, gender, race among many other criteria.
According to CDC (Central of Disease Control, USA), spirometry should be used as a precaution by the patients suffering covid

## Slide 3: Overview
Build a Orifice flow meter
Use Bernoulli's Equation to calculate Volume Flow Rate
Integrate Volume Flow Rate over time
differential Pressure Sensor
Transfer Voltage to Pressure Difference
Use machine Learning algorithm to Judge lungs' Performance.

## Slide 4: Project Demonstration
Bangladesh University of Engineering and Technology
Figure: Mouthpiece

## Slide 5: Orifice Meter
Bernoulli's Equation:
By applying this Equation for orifice meter:
Bangladesh University of Engineering and Technology

## Slide 6: Circuit Diagram
Components :
  Arduino uno
  LCD display
  Keypad
  MPXV5050DP sensor, A differential Pressure Sensor
Bangladesh University of Engineering and Technology
Proteus Simulation

## Slide 7: System Flow Chart:
Stop
Start
Input from user
Input from sensor
Calculation
Show result
Data from trained model
Bangladesh University of Engineering and Technology

## Slide 8: Calculation
Calculation of FVC from sensor data
V = Vp ×5/1024
Transfer equation: P=(V/5-.04)×1000/0.018
Q = (Ao ×((2 ×P / ρ)0.5)/((1-B4)0.5) ) ×106
Loop till Q is greater than a trigger value:
FVC=FVC+Q× S
Notation :
ρ= Density of air, 1.225 (kg/m3)
P= Differential pressure (Pa)
D = diameter of the pipe (m)
d = diameter of the orifice(m)
B = d/D
Vp= voltage output of sensor pin (v)
Ao = Area of orifice (m2)
FVC = Forced vital capacity (ml)
Q= Volume flow rate (ml3/s)
S=Data sampling time, 0.01s
Bangladesh University of Engineering and Technology

## Slide 9: Calculation: Calculation of Lung performance using Trained Model
Coefficients of trained model : M=
Data taken from user : X=
Lung performance :
  Y=MX'
[-67.497798082503 | 0.0237200392 | 0.617981660 | 29.6510904 | 2.23838912]
[ 1 | FVC | Age | Gender | Smoking status]
Bangladesh University of Engineering and Technology
Y(Lungs Performance) | Remarks
80% or greater | Normal
Less than 80% | Weak
Source: American Thoracic Society

## Slide 10: Lung Performance Dependency
Age
Physical Activities
Gender
Weight
Ethnicity
Height
Lung Functions Decreases Slightly each year with age
Lung Performance is Proportional to body size
Changes with different ranges of weight and height(BMI)
Plays a fundamental role, but not sufficient data to differentiate between all races
Lungs works better for people who exercise regular
Adult Female lungs is typically 10-12% smaller than males of same height and weight
Bangladesh University of Engineering and Technology
Source: Archivos de Bronconeumologia

## Slide 11: Data Availability
Test Subjects
Total Subjects: 1549
Age: Between 18 to 88
Train data: 0.9*1549 = 1394
Test data: 0.1*1549 = 155
Features
Gender
Age
Smoking Status
FVC(Forced Vital Capacity)
Male: 1224
Female: 325
Ex-smoker
Never smoked
Currently smokes
Bangladesh University of Engineering and Technology
Source: Kaggle

## Slide 12: Model Training (Language Used: Python)
Data Input
Preparing Data
Data Read
Selecting Needed Features
Data Cleaning
Conversion: String to Numerical Values
Linear Regression Model Training
Trained Matrix
Bangladesh University of Engineering and Technology

## Slide 13: Linear Regression Model
Least Square Method: The sum of the offsets or residuals of points from the plotted curve
Gradient Descent: An iterative method to analyze all weights with regard to small changes in error
Accuracy: 85-90%
Equation of Straight Line:
Y= mx + b
Multiple Linear Regression
y = b0 + m1b1 + m2b2 + ... mnbn
- Hyperplane
Bangladesh University of Engineering and Technology
Source: Toward Data Science

## Slide 14: Cost Analysis
Electrical Components | Price
Arduino Uno | TK. 700
10k Potentiometer | TK. 50
16x2 LCD Display | TK. 160
MPXV5050DP Sensor | TK. 2500
Jumper Wires | TK. 60
Matrix Key Pad (3x4) | TK. 90
Lodestar Soldering Iron | TK. 322
Solder Lead Mini | TK. 49
Total Cost | TK. 3931
Bangladesh University of Engineering and Technology
Source: Techshop bd

## Slide 15: Problems We Faced
The obstacles we encountered during the project:
1. Calibration.
2. Unavailability Of Components
3. Unavailability Of Sufficient Dataset : Total Patient, Features.
Bangladesh University of Engineering and Technology

## Slide 16: Calibration
Bangladesh University of Engineering and Technology
To accommodate, empirical discharge coefficient is used
The constant term K is evaluated by calibration.

## Slide 17: Unavailability of Components:
Due to the pandemic, we kept our search for hardware parts limited to online shopping.
We were able to find most of the electrical components. However, we could not find some hardware elements.

## Slide 18: Limitations
necessity of regular calibration
Spread out of germs/bacteria/virus
Bangladesh University of Engineering and Technology

## Slide 19
Thank You